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 New York City Police Department

PART I: Information and Requirements for Application 

For the Provision of Biotracks Expansion Project
PIN 056060000467
A.  Summary of the Solicitation  

The New York City Police Department (NYPD) is seeking one or more appropriately qualified laboratory that has, and maintains, accreditation by the American Society of Crime Laboratory Directors/Laboratory Accreditation Board  to provide requested DNA analyses of various evidence collected at crime scenes—hereinafter referred to as the Biotracks Expansion Project. 

Vendors interested, able and capable of provide the services required in the Biotracks Expansion Project should complete an application as further described in Part V below.  Submit the application to Frank Bello, Assistant Commissioner, NYPD-Contract Administration Unit, 51 Chambers Street, 3rd Floor, New York, N.Y. 10007 on or before ____________________.
B. Anticipated Contract Term

The NYPD anticipates that the contracts awarded from this solicitation will be for three years with two options to renew of two years each for a possible total of seven years. 
NYPD reserves the right prior to contract award, to negotiate alternate contract terms and renewal options, if any.

C. Application and Price Proposal Structure

Vendors interested in providing the above-described services should complete an application as described in PART  V, below. The application requests information to help the NYPD to determine applicant’s appropriate qualifications and experience. 
D. Contract Assumptions
1.
CONTRACT NEGOTIATIONS

a.
The NYPD will invite only applicants determined to be responsive to the requirements of this solicitation, and that are the most qualified to provide the services described herein, to enter into negotiations with NYPD.  Subsequent to initial negotiations, the NYPD may request that such  applicants submit a Best and Final Offer,  which the NYPD will evaluate pursuant to the same factors cited in Section E.4 below.
b.
The NYPD will request applicants, still under consideration, (upon completion of the evaluation processes and the Best and Final Offer processes, if any) to participate in further contract negotiations.  The NYPD will hold contract negotiations at the Police Department at a site to be determined.  All parties involved in the negotiation of City contracts will act in good faith in accordance with the City’s Charter and Procurement Policy Board Rules.

 2.
FAILURE TO NEGOTIATE

If a selected applicant fails to participate or provide the necessary information for negotiation in a timely manner or fails to negotiate in good faith, or cannot perform the contract within the amount of proposed, or within the budgeted funds available for the project, the City may terminate negotiations. 

3.
EQUAL ACCESS

Unless the NYPD deems that a question is of a proprietary nature to the inquiring applicant, the NYPD will send copies of  questions by prospective applicants and the answers thereto to all prospective applicants known to have received this solicitation document.  Only written responses from NYPD are official.

ADVANCE \d12
4.
BASIS FOR CONTRACT AWARD

The City will award contracts to responsible applicants whose applications are determined to be the most advantageous to the City, taking into consideration price* and the following factors listed in descending order of importance shall serve as the basis for determining the award of one or several contracts:

i. Experience providing DNA testing services and expert witness services. This category will be evaluated on the basis of the respondent’s reliability, testing accuracy, and customer satisfaction.  Respondent shall provide at least one of reference, but preferably more, which is currently using its laboratory service. In addition, Applicants shall provide at least one reference, but preferably more, of an agency collecting STR data for identification purposes.

ii. Approach and Methodology for conducting and completing the Tests and meeting the requirements in Part II. In addition to the applicants explanation of its scientific and laboratory approach and methodology used to conduct testing, the Committee will evaluate the responses description of its chain of custody or sample inventory system used to track samples from receipt in the contract laboratory to the return to the OCME and/or NYPD.  Additionally, the measures taken to ensure confidentiality of the kit samples and analyses will be evaluated. An on-site visit will be used to validate the respondent’s proposal and procedures. The on-site visit will be used to confirm use of the appropriate national standards and state protocols.

iii. Performance Standards and Quality Control. The Committee will evaluate the applicants:

a) Compliance with National Quality Assurance Standards;

b) Sample Handling and Security Protocols;

c) Quality Assurance Protocols;

d) Compatibility with OCME Databank Protocols; and

e) Appropriate Protocol for Troubleshooting.

f) Error rate, based on historical data, for testing DNA samples or kits, as described herein.

iv. Ability to Complete the Testing of Crime Kits in a Timely Manner. The City considers the services to be provided under this agreement to be time sensitive. Therefore, the respondent’s ability to get the job done in a timely manner will be an important consideration.

v. Price for the provision of the required, and or negotiated, services, inclusive of:

(i) Personnel and fringe benefits;

(ii) Shipping/transportation of kits and samples;

(iii) Management or other administrative overhead to the NYPD and OCME; and are the responsibility of the proposer and should be included in the price;

(iv) Overhead and Profit;

(v) Expert Witness Service Rates (for positive tests);

(vi) Performance Incentives and/or Disincentives.

(vii) Applicants shall provide a unit total cost for each kit and/or sample to provide the complete testing in accordance with the assumptions and the requirements (§II) stated herein.

(viii) Price proposals should be guided by the following:

· Payment made to Contractor(s) is expected to be performance based—that is dependent on the number of DNA tests performed without performance flaws, errors, malfunctions, and with minimal to no service downtime. 

· Applicants shall state in their Price proposal product and workmanship guarantees that satisfy performance-based assumptions. Applicants shall also state any recommended incentives for exceeding expectations, and disincentives for not meeting stated expectations. 

· Such terms shall be negotiated if needed.

Contract award will be subject to the timely completion of contract negotiations between the agency and the selected applicants.

5.
QUESTIONS

Questions regarding this  solicitation must be submitted in writing to the Agency Chief Contracting Officer at the address/fax/email below no later than ten (10) days after the Pre-Proposers Conference.

NYPD Agency Chief Contracting Officer

Contract Administration Unit

51 Chambers Street, 3rd Floor

New York, NY 10007

Fax # (646) 610-5129
Email:  fbello@nypd.org

6.
DOCUMENTS COMPRISING THIS SOLICITATION DOCUMENT

This Solicitation Document is comprised of the following documents:

a.  Solicitation Document, as amended, included all attachments and appendices.

b. APPENDIX A—STANDARD TERMS AND CONDITIONS AND SCHEDULE A WITH GENERAL CONDITIONS

c. APPENDIX B—ACKNOWLEDGEMENT OF ADDENDA

d. APPENDIX C—PRICE PROPOSAL FORMAT

Should there be any conflict among the documents; the following order of precedence will govern the resolution of conflict:

· First, the final negotiated agreement.

· Second, this Solicitation Document as amended or as modified by the NYPD's written answers to bidders' questions.

· Third, the Contractor’s technical and price proposal applications. 

· Last, the transcripts of the Contractor's oral presentation.
Part II—SCOPE OF WORK

ASSUMPTIONS AND REQUIREMENTS

A. 
 SCOPE OF WORK—ASSUMPTIONS AND REQUIREMENTS

Requirements and Specifications

 The Contractor shall perform services in accordance with the requirements and specifications  stated below and in the following attachments:

1.   The contractor shall participate in an external proficiency-test program from a test provider that has been approved by the American Society of Crime Laboratory Directors/Laboratory Accreditation Board (ASCLD/LAB). The contractor shall successfully complete proficiency tests throughout the term of this Contract. The Contractor shall provide a statement of compliance with same.

2.   The contractor shall ensure that each technician and analyst performing DNA analyses at the laboratory shall undergo open external proficiency tests semiannually.  To fulfill such criteria, the laboratory must ensure that one test be performed in the first six months of the calendar year and the second in the last six months of the calendar year.  In addition, the interval between consecutive tests must be at least four months and not to exceed 8 months.  Upon request, the contractor shall provide proficiency testing results for each analyst. 

3.   The laboratory shall maintain accreditation by the American Society of Crime Laboratory Directors/Laboratory Accreditation Board.  The Contractor shall provide a copy of a current accreditation certificate as well as copies of annual audits and/or inspections to the OCME Project Manager.  In addition, the Contractor shall obtain and maintain any applicable New York State DNA Testing Certification.

4.  The Contractor shall perform the requested DNA analyses to satisfy or exceed current standards for the quality assurance program for DNA analysis, issued by the Director to
the Federal Bureau of Investigation pursuant to the DNA Identification Act of 1994, entitled AQuality Assurance Standards for Forensic DNA Testing Laboratories;@ a copy of which is attached in ATTACHMENT 1.  The Contractor must follow all points in the discussion of the standards issued by SWGDAM.  The Contractor shall provide a certification that it shall perform the requested DNA analyses in accordance with the AQuality Assurance Standards for Forensic DNA Testing Laboratories.@

5.   The contractor shall process and analyze all samples in a secure, dedicated laboratory and in accordance with the Federal Bureau of Investigation’s NDIS standards for Acceptance of DNA Data and provide certification of such.

6.   The laboratory shall use either Applied Biosystems AmpF/STR Profiler Plus and AmpF/STR Cofiler kits, or Identifiler kits, as well as ABI analytical platforms (ABI 3730 or 3100 Genetic Analyer Instruments, or equivalent) and Windows NT based analysis tools.  Such equipment and tools are necessary in order to be consistent with the analyses performed by the OCME and to provide for complete integration with the existing operations of the local, state, and national DNA Databanks.  If technology improves, or if the market provides new technology or changes, the contractor must propose such changes to the OCME technical leader for approval before utilization of such new technologies or changes.  In addition, such changes must be consistent with OCME protocols and necessary validation studies must be completed.

7.   The Contractor shall supply all equipment and material necessary to comply with the OCME Protocols; a copy of which is attached as ATTACHMENT 2.

 8.   The Contractor shall keep on file, in its location, details and results of the laboratory’s process validation studies. Upon request of the NYPD or OCME, the Contractor shall provide documentation that demonstrates that it has carried out appropriate and adequate validation of its analytical procedures and data interpretation protocols.

9.  The transfer record (ie. chain of custody or LIMS records) for evidence and samples (generated from the evidence) shall document the inventory, processing, typing, and transmission of such items while in the custody of the Contractor as well as transfer of data, evidence, and samples back to the NYPD.  Any exchange of items between laboratory personnel shall require a written or documented transfer record.  Each relevant transfer record shall be included in each case file or control binder.

10.  The Contractor acknowledges that all work product generated as a result of this contract is confidential, and will treat such information with due care to prevent improper disclosure. Under New York State law, unauthorized disclosures of DNA records or the results of a forensic DNA test or analysis are proscribed as a Class E felony.  The NYPD reserves all rights to bring an appropriate civil action for any breech of confidentiality under the agreement.

11.  The Contractor shall protect the confidentiality of all records and other materials that are maintained in accordance with this contract. The Contractor shall have written policies governing access to, duplication and dissemination of all such information. The Contractor shall provide its employees and agents, if any, with a copy of a written explanation of these confidentiality requirements before access to confidential data is permitted. A copy of such written explanation signed by each employee shall be provided to the NYPD.

12.   The Contractor shall have and maintain the capacity to analyze the delivered cases withinthe specified time, herein.  In no instance shall a piece of evidence remain unanalyzed and/or completed by the Contractor more than sixty (60) days from receipt.  If case analysis is not completed by 60 days, the Contractor shall be reimbursed for 60% of case cost.

13.  Three of the most probative evidence items in a case shall be processed and analyzed.  If a reference sample is included, it shall be DNA typed.  If all three items in a case match the reference sample, additional items shall be tested with testing ending when a DNA type foreign to the reference sample is obtained.  Cases having more than 3 evidence items tested shall warrant a higher payscale, depending on the degree of additional testing completed (ie. Quantitation only vs. amplification and electrophoresis).  

14.  The Contractor shall describe and document all received evidence on worksheets (similar to those attached) and/or note paper (see ATTACHMENT 3).  

(a) Examination notes shall contain the individual case identifier, date of exam, examiner’s initials, description and measurements of said items.  



(b) The following items (or types of items) must be photographed (with scales in all photos)

 or sketched and tested areas must be clearly marked on such items and in such 



 photographs/sketches:




- clothing items




- weapons




- items that may have come in contact with individuals but do not have obvious




 stains (“used” tissues, shoes, scarves)




- any items with stained or unstained areas that will be further tested



(c) Written descriptions (or worksheets similar to ATTACHMENT 3) shall be used for the



 following items (photography or sketches are not needed for these and only these):




- beverage containers




- cigarette butts or other “smoked” items




- items “used by” individuals and have contact with their mouth (ie. cigarette




 butts, toothbrushes, lipstick/chapstick)




- swabs, crime scene samples



(d) The contractor shall perform presumptive and confirmatory testing for the following



 biological materials, as needed:




- presumptive testing for blood




- presumptive testing for semen,




- followed by p30 testing for confirmation



(e) All stains (biological or non-biological) and/or tested areas must be documented by



 diagrams and/or photography and must be given a unique identifying number, which shall



 be hand-marked by the analyst.  The contractor shall note the location of the stain, size,

 heaviness (soaked into fabric, surface smear, etc.).  Each photograph must have a ruler



 visible in the frame, either a plain straight ruler or an x, y axis ruler.



(f) Stained or unstained areas chosen for further testing shall be cut, swabbed, or scraped

and the chosen procedure shall be documented.  The OCME Department of Forensic



Biology “Scraping” procedure for the collection of biological material from fabrics has been

attached (ATTACHMENT 4).  This protocol shall be used to examine items where the identity of the “wearer” or “user” of such items is desired.  The Other Extraction Procedure is an optimal method to extract such “scraping” samples. (See ATTACHMENT 4)



(g) Samples identified as “control samples” by the NYPD do not need testing of any kind.



(h) If possible, the entirety of an item or sample should not be consumed during analysis.  It 



is recommended that at least 25% of the sample be saved for future analysis, if needed.  



However, if in the opinion of the analyst, consumption of the sample is necessary to have the 



best chance to obtain results, the item or sample may be consumed; the notes must clearly 



state this.

15.  Following extraction, a DNA quantitation shall be done using a method approved by the technical leader of the OCME Department of Forensic Biology.

16.   STR data shall be attempted for a minimum of 13 loci and Amelogenin. The following loci shall be used: 



D3S1358, D16S539, THO1, TPOX, CSF1PO, D7S820, VWA, FGA, D8S1179, D21S11,



D18S51, D5S818, and D13S317.

(a) Specimens shall be amplified as follows to fulfill duplicate amplification rules:

	 SEQ CHAPTER \h \r 1First Amplification
	Second Amplification (duplicate amplification)

	Identifiler
	Identifiler

	Cofiler
	Identifiler or Profiler Plus

	Profiler Plus
	Identifiler or Cofiler


    
(b) The analysis of a specimen shall not be considered complete until an attempt to genotype

 for all 13 loci and Amelogenin has been done.



(c) The genotypes shall be compiled in the appropriate format (Genotyper or Genemapper



 ID) for insertion into the FBI’s Combined DNA Index System (CODIS) and transmitted in



 electronic form on CD-Rom, while hard copies of electrophoretic data shall also be



provided in case specific files.



(d) All extraction, amplification, run worksheets and electropherogram data are to be



 returned to the OCME. All instrument data collection files and files generated in the



 analysis of the samples shall also be returned to the OCME on CD-Rom.  Control data shall



 be batched based on extraction sets and all subsequent controls paperwork generated



 regarding those extracts will be sorted together.  The batch numbers will be noted in the



 case files as well for ease in referencing between case file and control binders. 



(e) All unused sample portion and/or extracts are to be returned to the NYPD.  Amplified



 DNA shall be retained at the laboratory until appropriate quality assurance review has been



 completed by the OCME.  After such review, the amplified DNA shall be destroyed.



(f) Any quality control or quality assurance data, if not archived at the laboratory, is to be



 returned to the OCME.

17.   The quality control protocols used by the contractor shall be approved by the OCME.

18.   The Contractor shall monitor analytical procedures by using appropriate controls and                       standards on each run:



(a) 9947a is used as the positive amplification human DNA control.



(b) A negative amplification control and extraction negatives.



(c) In-lane size standard.

19.   The STR testing standards shall be in compliance with the following:



(a) All electrophoresis runs shall contain two Allelic Ladders.  All peaks on both Allelic



 Ladders shall be correctly labeled.



(b) All positive control (9947a) peaks shall be correctly labeled and identified.



(c) “No Size Data” shall be reported for the negative control and the reagent blank.



(d) No stutter peak shall exceed 20% of the peak height of the associated major peak.



(e) In each heterozygous category, the smaller labeled peak shall have a height of at least



 70% of the height of the larger labeled peak in the category.



(f) All peaks shall be labeled with base pair size rather with the designation AOL allele”.  In 



the event that an Off Ladder allele is observed, that allele must be confirmed by a separate 



electrophoretic run.  If the off ladder allele is confirmed to be an allele not present in the 



allelic ladder, this allele should be displayed as its base pair size.


(g) Any editing or analytical notes must be documented on a separate editing sheet in the



 case file (see ATTACHMENT 5).  The following are acceptable reasons for



 editing non-allelic peaks from a genotyper file:(From: Clark 1988, Walsh et al. 1996,


 
Clayton et al. 1998): 


! “Pull-ups” of peaks in one color caused by very high peaks in another color. This


 occurs only for multiplexes employing more than one labeling color and is caused by


the inability of the software to compensate for the spectral overlap between the


 different colors, if the peak height is too high. Pull-ups are matrix artifacts.


! Shoulder peaks approximately 1-4 bp smaller or larger than main allele (mostly


 larger), caused by a flat decline of main peak fluorescence.  Shoulders are easily


 recognized because they do not have the shape of an actual peak.


! -4 stutter peaks that are caused by slippage of the Taq polymerase enzyme during


 copying of the STR allele. Under rare circumstances there may be a +4 stutter peak. 


! Non-specific artifacts caused by non-specific priming in a multiplex reaction. These


 artifacts are usually easily recognized due to their low peak height and their position


 outside of the allele range.


! Labels placed on elevated or noisy baselines which do not resemble distinct peaks.


 Noisy and elevated baselines are matrix artifacts, otherwise known as spectral 


artifacts.


! “N” bands, where the main peak is split into two peaks caused by the Taq


 polymerase activity that causes the addition of a single A to the terminus of the


 amplified product (“N+1” band).  The allele calling is based on the N+1 bands,


 therefore complete extra A addition is desired. 


! Sharp spikes caused by power surges or crystals or air bubbles traveling by the laser


 detector window. Spikes might look like a single vertical line or a peak; they can


 easily be distinguished from DNA peaks by looking at the other fluorescent colors,


 including red; the spike is usually present in all colors.

(h) Plot containing labeled pull up peaks (artifactual signal due to signal in another color) may be accepted only under the following circumstances:


! The pull up peak shall have the same size in bp (+0.5 bp) as the corresponding off


 scale peak in the other color.


! Pull up peaks from no more than three of the 10 Profiler Plus categories, or from no


 more than two of the 7 Cofiler categories, may be disregarded in a given sample.

(i) Over-amplified or overloaded samples shall be addressed as follows.  Such samples often have peak heights between 5000 and 7000 fu’s and are characterized by a plateau shape or misshaped peaks, and often contain numerous labeled stutter peaks and artifacts. 



! In evidentiary samples, all DNA mixtures where peaks in at least one locus are 


6000 fu’s shall be rerun with less amp product.  All labels shall be removed from the



lane in question and list “numerous” for peaks removed on the editing sheet. 



 Exemplar or reference samples are exempt from this rule.  




! Overloaded samples that result in unacceptable pull up peaks, unacceptable split



 peaks, stutter bands or extraneous shoulder peaks, shall be re-run with less product.

20.   STR analyses of evidence samples shall undergo a 100% technical review by a second qualified analyst, using a technical review protocol approved by the OCME.

21.  All profiles generated through this contract shall be compared with (1) the DNA profiles of the contract laboratory’s staff and (2) all other generated profiles to identify potential contamination.

22.  Case files shall contain all relevant work product and data pertaining to the specific case and shall be assembled as follows:



! The left side of files shall contain chain of custody or LIMS paperwork, copies of



 vouchers, laboratory requests, and the case report.





! The right side of files shall contain (from the bottom up), (1) evidence examination



 notes and photograph, (2) in date order: Extraction paperwork (with batch numbers



 referenced), Amplification paperwork, Electropherogram data with editing sheet(s),



 and Quantification paperwork.

23.   Excess DNA extracts shall be stored frozen at -80C until such time that the Contractor can return the samples and/or extracts to the NYPD.

24.   The NYPD shall supply quality control (QC) samples with each shipment ( a minimum of 5% of each shipment).  If the results of the QC samples indicate a discrepancy from expected/known results, further submission and testing may be halted until resolution.

25.   The Contractor shall notify OCME and NYPD of any problems in testing, pertaining to this contract, with any corrective action measures immediately upon discovery (see ATTACHMENT 6).

26.   With respect to DNA analysis of the forensic casework samples, the Contractor agrees to provide expert witness testimony in subsequent legal proceeding, including but not limited to depositions, hearing and trials, as needed. The Contractor agree that it will provide certification that its cost proposal for travel fees, expert witness support, duplication of data and documents, and other litigation are equal to or less than those charged to other federal, state, local, or municipal law enforcement or district attorney clients. At the request of the NYPD, the Contractor shall provide copies of invoices documenting compliance with these requirements.

27.   Unless expressly approved by the NYPD, there shall be no subcontracting permitted under  this Agreement.

ATTACHMENT  TABLE OF  CONTENTS
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Attachment 5- Editing Sheet
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Part III-- Application Requirements

For the Provision of The Biotracks Expansion Project
A. Format and Content of the Application

Applicants shall provide one (1) original and six (6) copies of its proposal in the following format:

· A Cover Letter that introduces your company to the NYPD and that provides an overview of the DNA Testing services you are able to provide (limit this to two pages of text).

· In a Section Labeled “Section 1: Experience” describe your company’s experience in providing DNA testing services and expert witness services for Crime Enforcement purposes. Include an overview of the your corporate structure, including general mission statement, number of employees, list of current clients and other general information about the firm. Include a list of clients that you have provided such services to, with references and telephone numbers and the total amount of contract award for each. Provide information on the management team and staff that will be responsible for providing the services required herein. Explain your level of commitment to maintain your proposed staffing level and experience throughout the term of the agreement. Explain your ability to provide expert witness services and what such services will include. References will be factored into the evaluation. Failure to provide references that meet the requirements listed below will negatively impact the evaluation. Each laboratory shall provide three references from companies or public agencies not affiliated with the contractor.  At least one of these references, but preferably more, shall currently be using the laboratory service. At least one of the references shall be an agency collecting STR data for identification purposes.  Please provide the company agency name, address, a contact name, telephone number, facsimile number and E-mail address, if available. 
· In a section labeled “Section 2: Approach and Methodology” describe the scientific and laboratory processes your company will use to conduct and complete the DNA kit tests. Describe the facility(ies), equipment , resources, chain of custody approach. The respondent’s service proposal shall address the requirements listed in Section III of this Solicitation Document as explained in each part of that section. 
· In a section labeled “Section 3: Performance Standards and Quality Control” describe the product and workmanship guarantees that satisfy the NYPD performance-based requirements. In addition, Applicants shall provide information on it’s error rate, based on historical data, for testing DNA samples or kits, as described herein. 
· In a section labeled “Section 4: Ability to Complete Testing” include the dates and time frame that you can begin and complete the testing of crime kits.
· In a section labeled “Section 5: Price Proposal” Include the completed Appendix C Price Proposal Form included herein. Your price proposal shall be an all-inclusive itemized Price Proposal with no hidden or extra costs. If an item is not absolutely required it shall be listed as an optional item with a description of its functional usefulness and cost. Vendors shall also state any recommended incentives for exceeding expectations, and disincentives for not meeting stated expectations.
· Include as an Attachment to your proposal other information pertinent to the supply and services requested by the NYPD. Please limit your submission of other information to specifications and “cut-sheets” of products and services offered. 

B.  Due Date For Applications
Application packages are due on or before the date and the time, and at the location prescribed, as follows:_______________
Part IV - General Information To Applicants
A.   Complaints.  The New York City Comptroller is charged with the audit of contracts in New York City. Any proposer who believes that there has been unfairness, favoritism or impropriety in the proposal process shall inform the Comptroller, Office of Contract Administration, 1 Centre Street, Room 835, New York, NY 10007; the telephone number is (212) 669-3000.   In addition, the New York City Department of Investigation shall be informed of such complaints at its Investigations Division, 80 Maiden Lane, New York, NY 10038; the telephone number is (212) 825-5959.
B.   Applicable Laws.   This Solicitation Document and the resulting contract award(s), if any, unless otherwise stated, are subject to all applicable provisions of  New York State Law, the New York City Administrative Code, New York City Charter and New York City Procurement Policy Board (PPB) Rules. A copy of the PPB Rules may be obtained by contacting the PPB at (212) 788-7820.    


C.   General Contract Provisions.  Contracts shall be subject to New York City’s general contract provisions, in substantially the form that they appear in “Appendix A – General Provision Governing Contracts for Consultants, Professional and Technical Services".  A copy of the applicable document is either attached to the RFP or available through the Authorized Agency Contact Person. 

D.   Contract Award. Contract award is subject to each of the following applicable conditions and any others that may apply: New York City Fair Share Criteria; New York City MacBride Principles Law; New York City Burma Law; submission by the proposer of the New York City Department of Business Services/Division of Labor Services Employment Report and certification by that office; submission by the proposer of the requisite VENDEX Questionnaires/Affidavits of No Change and review of the information contained therein by the New York City Department of Investigation; all other required oversight approvals; applicable provisions of federal, state and local laws and executive orders requiring affirmative action and equal employment opportunity; and Section 6-108.1 of the New York City Administrative Code relating to the Local Based Enterprises program and its implementation rules.     
E.   Proposer Appeal Rights.  Pursuant to New York City’s Procurement Policy Board Rules, proposers have the right to appeal Agency non-responsiveness determinations and Agency non-responsibility determinations and to protest an Agency’s determination regarding the solicitation or award of a contract. 

F.   Multi-Year Contracts.  Multi-year contracts are subject to modification or cancellation if adequate funds are not appropriated to the Agency to support continuation of performance in any City fiscal year succeeding the first fiscal year and/or if the contractor’s performance is not satisfactory.  The Agency will notify the contractor as soon as is practicable that the funds are, or are not, available for the continuation of the multi-year contract for each succeeding City fiscal year.  In the event of cancellation, the contractor will be reimbursed for those costs, if any, which are so provided for in the contract.
G.   Prompt Payment Policy.  Pursuant to the New York City’s Procurement Policy Board Rules, it is the policy of the City to process contract payments efficiently and expeditiously.  
H.   Prices Irrevocable.   Prices proposed by the proposer shall be irrevocable until contract award, unless the proposal is withdrawn.  Proposals may only be withdrawn by submitting a written request to the Agency prior to contract award but after the expiration of  90 days after the opening of proposals.  This shall not limit the discretion of the Agency to request proposers to revise proposed prices through the submission of best and final offers and/or the conduct of negotiations.

I.   Confidential, Proprietary Information or Trade Secrets.  Proposers shall give specific attention to the identification of those portions of their proposals that they deem to be confidential, proprietary information or trade secrets and provide any justification of why such materials, upon request, shall not be disclosed by the City.  Such information must be easily separable from the non-confidential sections of the proposal. All information not so identified may be disclosed by the City.

J.  Solicitation Document Postponement/Cancellation.  The Agency reserves the right to postpone or cancel this SOLICITATION DOCUMENT, in whole or in part, and to reject all proposals.  

K.  Proposer Costs.  Proposers will not be reimbursed for any costs incurred to prepare proposals.


L.  Charter Section 312(a) Certification. 
The Agency has determined that the contract(s) to be awarded through this Solicitation Document will not directly result in the displacement of any New York City employee.

___________________________________

___________________

           Agency Chief Contracting Officer


 Date

APPENDIX B- ACKNOWLEDGMENT OF ADDENDA

PIN 056060000467— Provision of DNA BIOTRACKS PROJECT
Instructions:
The respondent is to complete Part I or Part II of this form, whichever is applicable, and sign and date this form. This form serves as the respondent’s acknowledgment of the receipt of the Addenda to this Solicitation Document which may have been issued by the NYPD prior to the Proposal Due Date and Time.


Part I:  Check Box if Applicable:

Listed below are the dates of issue for each Addendum received concerning this Solicitation Document:

Addendum # 1,  dated:   ____/____/____
Addendum # 2, dated: ____/____/____

Addendum # 3,  dated:   ____/____/____
Addendum # 4, dated: ____/____/____

Addendum # 5,  dated:   ____/____/____
Addendum # 6, dated: ____/____/____

Addendum # 7,  dated:   ____/____/____
Addendum # 8, dated: ____/____/____

Addendum # 9,  dated:   ____/____/____
Addendum # 10, dated: ____/____/____

Addendum # 11,  dated:   ____/____/____
Addendum # 12, dated: ____/____/____

______________________________________________________________________________

Part II:
 Check Box if Applicable:   

No addendum was received in connection with this Solicitation Document.

Respondent’s Company Name: ____________________________________________________

Respondent’s Authorized Representative:


Name:__________________________________________________________


Title:___________________________________________________________


Signature:_______________________________________________________


Date:___________________________________________________________
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